**Prada Lab Homogenizer Care**

Because this is a metal instrument used in salt water, the following steps are crucial to preventing the motor from rusting. All salt must be removed after use and careful use is required to prevent any water from going up into the motor.

How to use homogenizer

* Clean homogenizer before and after each use
* **Never submerge probe more than 2 inches**
* **Never lay homogenizer down horizontally**

How to clean and store homogenizer

* Run the homogenizer in DI for 30 seconds
* Unplug
* Remove the probe cover and store next to stand
* Use DI squeeze bottle to rinse off any areas of the probe that were exposed to salt water
* Use a paper towel to pat the probe dry
* Store on stand so that no liquid can move up into motor